| PI: Tyler Doerksen Contact: | Phone 928-925-8386 |  | INVOICE NO. TSD3112251A |  |
| --- | --- | --- | --- | --- |
| Company/University & Bldg w/ Room# |  | Date of Service 12/31/2025 |  | Next Service Due December '26 |
| # Balance Information ## Model: XS105DU Serial No: B621501811 Balance Cal Due Date: 31 August 2026 |  |  |  |  |

CALIBRATIONS

INTERNATIONAL INC.

Toll Free 1-877-CAL-INTL

(1-877-225-4685)

2142 Chipset Street  Bozeman, Montana 59718

theresa@calibrationsinternational.com

406-589-6025	BASIC CALIBRATION SERVICE

| PIPET NO. |  | MANUFACTURER/ MODEL | MAX VOL. | IN SPEC | OUT SPEC | ADJUSTMENTS & REPAIRS* |  |  |  |  |  |  | FINAL CORRECTED VOLUMES |  |  |  | P/F |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | SE | COR | G | SH | FR | O | OTHER | SAMPLE 1 | SAMPLE 2 | SAMPLE 3 | SAMPLE 4 |  |
| 01 | C202358375 | Biotix Xpipette | 200 | * |  | no seal |  | K |  |  |  |  | 199.9 | 199.5 | 199.6 | 199.7 | P |
| 02 | F03069 | Rainin Pipette Lite | 200 | * |  | * |  | K |  |  |  |  | 99.5 | 99.6 | 99.9 | 99.8 | P |
| 03 | B0964962A | Rainin Pipette Lite | 200 |  | *high | X |  | K |  |  |  |  | 9.95 | 9.98 | 9.99 | 9.96 | P |
| 04 | F0104131A | Rainin Pipette Lite | 20 | * |  | * |  | K |  |  |  |  | 199.5 | 199.6 | 199.8 | 199.7 | P |
| 05 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 06 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 07 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 08 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 09 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 16 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 17 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 18 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 19 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

*SE(Replace Seal), COR(Corrosion), G(Grease), SH(Shaft), FR(Friction Ring), O(O Ring)

An Internal Q.C. Calibration check was performed by CALINTL Technician on Balance noted, prior to calibration of pipettes on this date.

Balance Calibration Certificate indicating NIST Traceable available.  CALINTL Technician calibrated pipettes to meet ISO 8655-2:2002

Notes: _________________________________________________________________________________________________________________________

Verified By:  CUSTOMER EMAIL:  tyler.tyler.d@gmail.com   PAGE     1     OF    1

Group:  __________   Group Coordinator:  ________________________________  Coordinator Email:  _________________________________

CALINTL BASIC SERVICE 10/19